| Draft |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| ATLAS COOLING OF THE DETECTOR TEST IN MAGNETIC fIELD |  |  |  |  |  |  |  |  |
| Objectives To validate material in UX cavern Abstract Test of a piezzo pressure transmitter in a magnetic field. Click here to access to the guidelines |  |  |  |  |  |  |  |  |
| Prepared |  |  |  | Verified |  |  | Approved |  |
| By: | BONNEAU Pierre |  |  | By: | Reviewer |  | By: | Approver |
| Date: | 2002/09/24 |  |  | Date: | 200Y/MM/DD |  | Date: | 200Y/MM/DD |
| Distribution List |  |  |  |  |  |  |  |  |
| History of Changes |  |  |  |  |  |  |  |  |
| Rev. No. |  | Date | Pages |  |  | Description of Changes |  |  |
| 1 |  | 2002/10/07 |  |  |  | Add 2nd test up to 1.5[T] |  |  |

Table of contents

1.	Description	3

2.	Results	4

3.	second test	6

4.	conclusion	9

# Description

We have tested a pressure transmitter in a magnetic field. The transmitter is an HUBA OEM type 501 with ceramic piezzo technology, pressure range from 0 to 10 bar relative, output 0-10V, (Figure 1). Exact serial number is 501.930002111 and detailed technical characteristics are in annex: Press_501e.pdf .

Figure 1 – HUBA Pressure transmitter OEM type 501

The transmitter is connected to nitrogen supply bottle through a pressure regulator to have a variable pressure. We have a reference pressure transmitter (HUBA 680.71121100, 0/16 bar.a) connected in parallel outside the magnetic field. Both sensors are connected to a Labview DAQ system. The 501 is connected with a 5m shielding cable (Figure 2).

We install the transmitter in horizontal and vertical position inside the magnet MNP17 in the building 168 – Thanks to R. Rey-Mermier and Felix Bergsma of EP-TA3. This magnet allows field from 0 to 2700 [Gauss] with presetting steps of 85, 171, 342, 684, 1350 and 2700 [Gauss].

Figure 2 – Layout of the test

We tested the behaviour of the transmitter in the magnetic field in two ways:

- We applied a constant pressure of 0, 5 and 10 bar and increased the magnetic field from 0 to 2700 [Gauss] on each fixed pressure.

- We applied a constant magnetic field of 0, 684, 1350 and 2700 [Gauss] and increase manually the pressure from 0 to 10 bar.

We proceeded to these 2 measurements with the sensor in horizontal position and repeated the same with the sensor in vertical position.

# Results

We don’t see any effect of the magnetic field on the pressure transmitter in any of the 4 tests. The curve of the tested sensor follows the curve of the reference sensor.

Figure 3

Figure 4

Figure 5

Figure 6

# second test

We will have transmitters located at the outer edge of the calorimeters to control the pressure of the cooling system for TILE and LAr detectors. The magnetic field in this region is expected to be 1 to 1.5 [T] (or 15000 [Gauss]), so we made a second series of test on the MEP-24 of EP-TA3, that gives this level of magnetic field.

We proceeded at the same 4 tests with setting point of 0.4, 0.8, 1.2 and 1.5 [Tesla].

No effect in the transmitted signal can be observed.

Figure 7

Figure 8

Figure 9

Figure 10

# conclusion

This low cost OEM sensor (around 70 CHF for 400 pieces) proposed by the Huba Control Company is magnetic resistant up to 1.5 [T] and can be used for the control of the cooling systems in the cavern UX.